## Analysis method
Appendix 3. Analysis method |  | Appendix 3. Green Procurement Standards_Ver. 9.1
rev.01
(1) Analysis method conform to IEC62321.
(2) A quantitative analysis must be conducted on each subpart of the item. The result must be within the threshold level.
(3)If the combination of pretreatment and analyzers can guarantee the following thresholds for the lower limit of quantification at the time of analysis,
analysis methods other than those shown in the examples are also acceptable.
Less than 5ppm of mercury, less than 5ppm of cadmium, less than 2ppm of total chromium, less than 30ppm of lead.
(4) If the test data are to be used as an analysis certificate, describe the test method used.
(5) No analysis of PBB and PBDE is required by this Standards document, however, this method is recommended for analysis if it is to be conducted to meet
Customer's requirements.
(6) No analysis of DEHP, BBP, DBP and DIBP is required by this Standards document, however, this method is recommended for analysis if it is to be conducted
to meet Customer's requirements.
Substance name | Preliminary processing: | Analytical instruments:
Cadmium/Cadmium compounds, Lead/Lead compounds, Chromium/Chromium compounds | Prepare a solution by wet decomposition method (including pressure decomposition) in the presence of, for example, sulfuric acid, nitric acid, hydrochloric acid, hydrofluoric acid, or hydrogen peroxide; by incineration in the presence of sulfuric acid; or by pressurized acid decomposition (microwave decomposition). If precipitation occurs in the solution, dissolve it completely into a solution by, for example, hydrofluoric acid decomposition or by alkali fusion decomposition. | Examples - Inductively Coupled Plasma Atomic Emission Spectrometer (ICP-AES [ICP-OES]), inductively coupled plasma mass spectrometer (ICP-MS), atomic absorption spectrometer (AAS, FLAAS).
Mercury/Mercury compounds | Dissolve into a solution using, for example, sulfuric acid, nitric acid, or potassium permanganate; and in a pressure decomposition flask, or a decomposing flask equipped with a reflux condenser to prevent mercury emission. A microwave decomposition method may also be used. | Examples—Reduction vaporization ICP emitted light spectrometer (ICP-AES [ICP-OES]), atomic absorption spectrometer (reduction vaporization, heated vaporization), inductively coupled plasma mass spectrometer (ICP-MS).
Hexavalent chromium compounds | Elute Hexavalent chromium into water solution by, for example, by hot water extraction, or decomposition in alkali solution, etc. | Example—absorptiometer and ion chromatograph, etc.
PBB and PBDE (for reference only) | Dissolve in an organic solvent or extract with a Soxhlet extractor. Use the silica gel column process to obtain a solution. | Gas chromatograph mass spectrometer (GC-MS)
DEHP, BBP, DBP and DIBP (for reference only) | Analysis method conform to IEC62321-8. Determination of certain substances in electrotechnical products - Part 8 | Gas chromatograph mass spectrometer (GC-MS)